## Slide 1: Rosetta Magnetic Field PDS Review
B. J. Anderson

## Slide 2: Completeness of data
Science needs: See actions 1, 8, 10-11.
Calibration levels: See actions 2-3, 5-7, 10-11.
Housekeeping and science operations info.: See action 4.
Software: Not evaluated.
Documentation: See actions 9, 12-13.
Calibration information: See actions 10, 12-13.

## Slide 3: PDS Compliance
Data structures: appear to be PDS standard.
Conventions used: standard time and magnetic field conventions used.
Geometry information: Did not see SPICE kernel information – if these are provided under another Rosetta delivery they should be referenced in the MAG documentation.
SIS document: Seems complete.
Completeness:
  Documentation: YES (with few minor corrections)
  Catalogs: Not evaluated.
  Index files: Not evaluated.
  Processing levels: YES.
  Data coverage: YES.

## Slide 4: Recommended Revisions (1-3 of 12)
Time offsets due to PIU (Power or Plasma Interface Unit?) software filter algorithm can be large (>100s). This is not applied to OBT but only to UTC in the products. Action: Need to clarify that this time correction was accounted for when relating the heater, reaction wheel, and other spacecraft events (power system transitions etc. – though not mentioned) to magnetic field signals in the IB and OB sensor data.
Corrections for Lander Heater Current, and Reaction Wheel signals were applied. Action: Spacecraft events (heater currents, solar array angles, high-gain antenna angles, reaction wheel rates) need to be provided (are not included in delivery) and registered in UTC since the PIU time offset is only included in the UTC magnetic field data.
Temperature corrections were critical in calibrating the data. Action: Need to clarify the following: Was the time offset applied in doing the in-flight temperature correction calibration? Or were the temperature and magnetic field data aligned in OBT? (Which would be wrong according to the long time offsets that the PIU software introduces.)

## Slide 5: Recommended Revisions (4-7 of 12)
Spacecraft (SC) currents evidently contribute to the magnetic field signals. Action: Specify where SC telemetry for currents and gimbaled structures are provided in PDS delivery. If not in another data set then provision for them should be made in the magnetometer delivery.
Correction for reaction wheel signals is discussed in the frequency domain. Action: Need to clarify the following: How was the phase of the reaction wheel signal established? How was the conversion back to the time domain done (presumably via inverse FFT). What time windows were used for the FFT and IFFT to implement the correction (i.e. the frequency resolution of the reaction wheels)?
Data qualify flag is very useful. Action: The definitions of perfect, good and poor correlation between IB and OB sensors are not given.
Clarification of data qualify flag definition: Action: The measure used to determine whether IB and OB show different long-term behavior is not given. That is, what does long-term mean? Longer than what time interval?

## Slide 6: Recommended Revisions (8 of 13)
The quality flag string is very useful. It would perhaps be useful to indicate the quality flags which would be considered appropriate for science analysis. Action: Add set of flags that can be regarded as valid science data.
Presumably these are:
Char 1: 0 or 1
Char 2: 0 or 1
Char 3: 1 (i.e. one would assume that only boom deployed data are appropriate for science).
Char 4: 0
Char 5: 0 or 1 (*see notes below)
Char 6: 0
Char 7: x
Char 8: x
Thus, the following eight strings indicate valid science data:
“xx000100”, “xx010100”
“xx000110”, “xx010110”
“xx000101”, “xx010101”
“xx000111”, “xx010111”
Is this true?

## Slide 7: Recommended Revisions (9-11 of 13)
The data set under review consists of only a few days of data. Action: Would it not be reasonable to request that some form of survey overview plots generated by the team be provided in a platform independent format, e.g. png or jpeg? (Note that such plots are included in flight reports.) Separate stand-alone plot files would be extremely useful to investigators trouble shooting their own implementations of file readers and displays.
Extensive information on the amplitude response of the magnetometers with frequency are provided. However, no information about the corresponding phase relationship is given. Action: The calibrated phase function vs. frequency and/or fixed time delay in the instrument should be given.
No information is given about the PIU filter and its frequency response or its phase response. Action: Since the PIU time offset is dramatic, the frequency and phase responses of the PIU software should be documented as well.

## Slide 8: Recommended Revisions (12-13 of 13)
Calibration gain, temperature, offset etc. A veritable mountain of data and plots are provided but it is not until page 163(!) of the RO_IGM_TR00003.pdf document that a useful summary of the results is presented. This should be given in a suitably named summary of calibration results document so that it can be readily located. In this regard, RO_IGEP_TR0028.pdf is a very useful document. Action: The calibration (RO_IGEP_TR0028.pdf) summary should be suitably named so that investigators can readily find this information.
The poor quality of the figures in RO_IGEP_TR013.pdf (and .doc) is unfortunate. These are highly compressed graphic files, which are barely legible. Action: Perhaps better quality originals are still available?