## Slide 1: R-Alice PDS/PSA Data Review PRL & ESC1 phases
Cesare Grava

## Slide 2: Datasets Reviewed
RO-C/CAL-ALICE-2-PRL-V1.0
RO-C/CAL-ALICE-3-PRL-V1.0
RO-C/CAL-ALICE-4-PRL-V1.0
RO-C/CAL-ALICE-2-ESC1-V1.0
RO-C/CAL-ALICE-3-ESC1-V1.0
RO-C/CAL-ALICE-4-ESC1-V1.0
Review Conclusions
  Data are ready to be released after some RIDs are addressed
  RIDs submitted: 101,102,103,104,105

## Slide 3: R-Alice Spec’s
Spectral Range 700 - 2050 Angstroms
Spectral Resolution 12.5 - 9.8 Angstroms for extended sources
Spatial Resolution 0.6°
Projected Entrance Slit 0.1° by 6°
Nominal Effective Area 0.03cm2 (at 1900 Angstroms), to 0.53cm2 (at 1150 Angstroms)
Very similar to the LRO/LAMP UV spectrograph I work with.

## Slide 4: Evaluation Tools
Laptop Lenovo Thinkpad:
  Virtual Machine Oracle Fedora 17x86_64
  Windows 7
Data installed just fine with tar; had to download multiple times the huge tgz files though…
IDL 8.4
Readpds v4.8
Astron library

## Slide 5: RO-C/CAL-ALICE-2/3/4-PRL/ESC1-V1.0
VOLDESC.CAT, and ERRATA.TXT: good
AAREADME.TXT: see RID 103

## Slide 6: RO-C/CAL-ALICE-2/3/4-PREL/ESC1-V1.0
CALIB/CALINFO.TXT: good
CALIB/RA_AEFF_008 and its corrections (see RIDs # 101, 102):
  RA_CR2_PC10_CORRECTION
  RA_CVP2_PC10_CORRECTION
  RA_ESB2_PC10_CORRECTION
  RA_ESB3_PC10_CORRECTION
  RA_MARS_PC10_CORRECTION
  RA_PC4_PC10_CORRECTION
  RA_PC6_PC10_CORRECTION
  RA_PC8_PC10_CORRECTION
  RA_STEINS_PC10_CORRECTION
CALIB/RA_PC12_AEFF_001 and its correction:
  RA_LUTETIA_PC12_CORRECTION
CALIB/Flat, Dark, and Wave (2,3,4): good

## Slide 7: RO-C/CAL-ALICE-2/3/4-PREL/ESC1-V1.0
CALIB/RA_DARK_3900V_V003.FIT

## Slide 8: Effective Area Calibration Data
CALIB/RA_AEFF_008: read with readcol.pro, no problem

## Slide 9: CATALOG
CATALOG/ALICE_RO.CAT  SOFTWARE.CAT (2,3,4): good
CATALOG/REF.CAT: good. RID 105: Could add some more references (Stern et al. 2015; Feldman et al. 2015; Feaga et al. 2015)

## Slide 10: DOCUMENT & INDEX
DOCUMENT/* (2,3,4): good
  RID 104: change name of ALICE_L3_TO_RAYLEIGHS.ASC
  DOCUMENT/ALICE_DATA_TO_RAYLEIGHS.ASC is especially nice to have.
    Note that conversion from flux photons/cm2/sec to differential flux (i.e., specific intensity) photons/cm2/sec/A by dividing by Δλ/pix is described here with an idl code example. I tested it and it did work!
DOCUMENT/CODE/* (2,3,4), Entire “MIKE” calibration pipeline code: good
  Didn’t get around to running myself, but this is a nice addition since it explains how the calibration is done (most users won’t ever need this).
INDEX/ (2,3,4): good
  2 is ENG, uncalibrated
  3 is SCI, flux-calibrated
  4 is LIN, extracted calibrated spectra

## Slide 11: DATA
Naming convention:
  <TYPE> is PIX# (pixel list), HIS# (histogram mode), CNT# (count rate)
    and where # is the 1-digit version number to distinguish different processing runs of the data (i.e. different improved calibrations);
  <STD> is ENG, SCI, or LIN level (CODMAC 2, 3, or 4, respectively);
Read data both with rdfits_struct.pro (for .fit) and readpds.pro (for .lbl): no problem.
RA_<YYMMDDhhmmss>_<TYPE>_<STD>.<EXT>

## Slide 12: DATA
Successfully plotted SCI and LIN versions of one ESC1 spectrum of 67/P. I applied the instructions (from alice_data_to_rayleighs.asc) for converting SCI to LIN and I got (correctly) the LIN spectrum

## Slide 13: DATA
Reproduced Fig. 2 (right) and fig. 3 (left) of Feldman et al. 2015
DATA

## Slide 14: Summary
Data are complete, intelligible, interpretable, and of archival quality
RIDs: two major, one minor, two editorial
  101,102,103,104,105
  Listed on the next five slides

## Slide 15: RID 101
RID Classification: Major
Location: calibration/
Dataset : ALL
Title: number of elements in Effective Area file
Description: The effective area, with 668 elements, does not correspond in size with the number of used columns in the EDR/SCI/LIN files (1024 – 100 = 924, as explained in the EAICD). There are obviously other columns that are unused.
Recommended Solution: It should be described somewhere (perhaps in the EAICD) the exact range of the effective columns of the illuminated detector.

## Slide 16: RID 102
RID Classification: Major
Location: calibration/
Dataset: ALL
Title: number of elements in Effective Area correction files
Description: The effective area file, with 668 elements, does not correspond in size with the series of “effective area corrections” files, which have 1024 elements.
Recommended Solution: It should be described somewhere (perhaps in the EAICD) how these “effective area corrections” are applied.

## Slide 17: RID 103
RID Classification: Minor
Location: AAREADME.TXT
Dataset: ALL
Title: Missing mention of LIN dataset
Description: in the file aaareadme.txt there is no mention of the third data product of Alice, probably the most used one (LIN): the “linearized”, resampled, calibrated data (CODMAC level 4).
Recommended Solution: Add the description of the LIN data products after SCI data products.

## Slide 18: RID 104
RID Classification: Editorial
Location: docinfo.txt
Dataset: ALL
Title: change name of ALICE_L3_TO_RAYLEIGHS.ASC
Description: In docinfo.txt, the document referred to as ALICE_L3_TO_RAYLEIGHS.ASC is in reality ALICE_DATA_TO_RAYLEIGHS.ASC
Recommended Solution: Replace in every /Document/docinfo.txt ALICE_L3_TO_RAYLEIGHS.ASC with ALICE_DATA_TO_RAYLEIGHS.ASC.

## Slide 19: RID 105
RID Classification: Editorial
Location: ref.cat
Dataset: ALL
Title: update reference list
Description: three recent references could be added to the list
Recommended Solution: Stern et al. 2015, Feldman et al. 2015, Feaga et al. 2015 can be added to the list.