## Slide 1: New Horizons Alice PDS-SBN Review
Lori Feaga
May 19, 2016

## Slide 2: NH Alice Data
Overall – in good shape

## Slide 3: Data Set Inconsistencies
aareadme.txt – (and aareadme_bu.txt files in both Level 2 & 3) incomplete thought and sentence in the Suggested FITS & PDS Software section
voldesc.cat – (Level 2), which includes a calib directory, should include calibration information in its DESCRIPTION
Spectrograph vs. spectrometer terminology!

## Slide 4: calinfo.txt (Levels 2 & 3)

## Slide 5: Calibration Files Used in Pipeline (for a subset of data)

## Slide 6: Data Set Typos
See Lori Feaga’s submitted .txt file with full listing of typos

## Slide 7: Alice Slit Shape
Row 19 (zero indexed)
Row 6 (zero indexed)
Row 25 (zero indexed)
Row 18 (zero indexed)
Solar
Occultation
Channel

## Slide 8: Lollipop “fuzz”
To reduce sensitivity at the bright HI Lyman-α emissions at 1215 Å, a 70 Å band of the photocathodes were not coated with KBr or CsI. However, in the 2°x2° box portion of the slit, the HI Lyman-α spreads out beyond the uncoated portion of the detector an additional ~110 Å. So, for ~90 Å [(70 + 110 )/2] on either side of the coated photocathodes, there is contamination from HI Lyman-α that has been dubbed the lollipop fuzz.
Wording in the dataset.cat:
“However, in the 2x2 degree box portion of the aperture designed to capture the Sun during occultations, the H Lya spreads out beyond the uncoated 70-Angstrom band over another ~110 Angstroms of photocathode-coated, more sensitive detector on either side.”
λ 

## Slide 9: Lollipop Fuzz ~180 Å Wide Verified Spectrally

## Slide 10: Basic Data Analysis

## Slide 11
Checked data with NASAVIEW, SAOImage ds9, IDL ReadPDS
  could display data with labels as expected

## Slide 12: SPICE Geometry Check
Sunward Pointing (2015-06-16T16:41:38)
  PDS label shows Sun to target distance of 0 km ✓
  .fit file header agrees with SPICE for sub-S/C lat and long on Sun (5°,33°) ✓
Pluto Pointing (2015-07-13T06:41:40)
  PDS label agrees with SPICE for S/C to target distance of 1444727 km ✓
  .fit file header agrees with SPICE for boresight RA & dec, sub-S/C lat and long, etc. ✓

## Slide 13: Increase in SNR w/ Exposure Time
1 sec
10 sec
100 sec
1000 sec

## Slide 14: Scaled Sun Overplotted
Sun doesn’t match the spectra –
but should only be in one pixel,
so Sun not in the narrow slit

## Slide 15: Scaled Sun Overplotted
Sun matches the spectra –
Sun in the wide slit (SOC)

## Slide 16: Pluto

## Slide 17: Pluto
Rows 7,8
Rows 19,21
Row 20

## Slide 18: The End