## Slide 1: SDC PDS Review Pluto Data: January 2015 to January 2016
A. R. Poppe
U.C. Berkeley
December 5, 2016

## Slide 2: SDC Instrument
SDC instrument:
- 14 detectors: 12 science (with 1 defunct), 2 reference
- electronics box for signal processing, telemetry/commands

## Slide 3: Previous & Current Releases
SDC launch, jupiter, and cruise datasets cover 1/2006 to 7/2014, Earth to approximately the orbit of Neptune
Three main data periods previously released by SDC:
Launch: March 2006 – Dec 2006, mainly testing and checkout of instrument
Jupiter: Dec 2006 – June 2007, small chunk of data taken post-Jupiter fly-by (note SDC was not on during close approach of Jupiter)
Pluto-cruise: June 2007 – July 2014, main period of scientific data collection by SDC between the orbits of Jupiter and Pluto
Current dataset:
Pluto Encounter: Jan 2015 – Jan 2016, Pluto approach, encounter, post-encounter

## Slide 4: SDC Raw Data
Includes SDC raw scientific data, SDC housekeeping (engineering telemetry), New Horizons spacecraft housekeeping (0x004, 0x00D, 0x00A), and New Horizons thruster firing fata
The SDC raw scientific data includes channel (detector) number, data threshold, data magnitude, and time stamp

## Slide 5
Plot of Mission Elapsed Time vs entry looks good
Total number of events: 10,118
Time period covers: Jan 17, 2015 – Jan 23, 2016
- note: this is about 6 months longer than the data reviewed at the May ‘16 meeting
SDC Raw Data

## Slide 6
Channel ID# versus Time looks good
Total number of events: 10,118
Time period covers: Jan 17, 2015 – Jan 23, 2016
- note: this is about 6 months longer than the data reviewed at the May ‘16 meeting
SDC Raw Data

## Slide 7
Magnitude versus Time looks good – note threshold change post Pluto encounter
Total number of events: 10,118
Time period covers: Jan 17, 2015 – Jan 23, 2016
- note: this is about 6 months longer than the data reviewed at the May ‘16 meeting
SDC Raw Data

## Slide 8: SDC Calibrated Data
Includes SDC calibrated scientific data
The SDC calibrated scientific data includes time, channel, charge, mass, mass threshold, error bars, quality flag, and theoretical SDC-dust relative impact velocity

## Slide 9: SDC Calibrated Data - ERROR
For some reason, these two files have different dimensions (9 cols) than the properly-formatted SDC science files (10 cols)…
These two files are incorrectly formatted:
nh-p-sdc-3-pluto-v2.0/data/20160119_031574/sdc_0315748362_0x700_sci.fit
nh-p-sdc-3-pluto-v2.0/data/20160121_031574/sdc_0315748363_0x700_sci.fit
This is the FITS header for a correctly formatted level 3 file:

## Slide 10
Plot of Mission Elapsed Time vs entry looks good
Again, note failure of Channel 11 (described in documentation)
Significantly increased count rate post-Pluto encounter… due to a return to nominal, lower thresholds
SDC Calibrated Data

## Slide 11
SDC minimum thresholds were dropped by about 1 order-of-magnitude post Pluto encounter
This leads to a greater number of events.
SDC Calibrated Data

## Slide 12
SDC minimum thresholds were dropped by about 1 order-of-magnitude post Pluto encounter
This leads to a greater number of events.
SDC Calibrated Data

## Slide 13
Incorrect on/off table for SDC has been successfully resolved from May 2016 review
SDC Calibrated Data: On/Off

## Slide 14: Textual Edits: nh-p-sdc-[2,3]-pluto-v2
Aareadme.txt:
  Under “Whom to contact…”: Mihaly’s name is mis-spelled, and is duplicated two lines in a row
Dataset.cat, Data Set Overview,1st paragraph:
  “New Horizon” -> “New Horizons”
Dataset.cat, Data Set Overview,3rd paragraph:
  “...limit is lower than earlier in the mission...”
Document/docinfo.txt:
  LUNINEETAL1995.PDF: “background” is mis-spelled
Document/nh_mission_trajectory.lbl:
  DESCRIPTION: “instrument” mis-spelled
Data/*/*.lbl:
  Under “QUALITY_FLAG” description, “initials” is mis-spelled

## Slide 15: Misc questions - nh-p-sdc-[2,3]-pluto-v2:
document/sdc_seq_pluto.tab
  There’s only entry here… is this correct?
  Differs, for example, from the plutocruise dataset, which has several sequences listed

## Slide 16
Completeness:
  Each time period (launch, jupiter, cruise, pluto) has a complete set of files and documentation
  Adequate completeness of each data record (i.e., one dust hit)
Intelligibility:
  Documentation is adequate and straightforward
  SDC data are intelligibly formatted through FITS files, both raw and calibrated
  Metadata are straightforward through ASCII implementation
Interpretability:
  Given the relatively simple nature of SDC operations, the data are straightforwardly interpretable; no issues noted here
Only one major lien:
  Need to correct two files in the nh-p-sdc-3-pluto-v2.0 dataset