## Slide 1: Review of the LEISA Data Set
Karl Hibbitts
June 19, 2017

## Slide 2: Calibrated V3 data
FITS files open easily and data extracted via Matlab
  Data files
  Metadata files
AAReadme: clear information
AARINFO: clear information
SOC Instrument ICD: Oct 2016; Useful and accurate
  54 high res channels
  198 nominal res channels
  Details are left to the user to determine

## Slide 3: Review of Metadata Files
Wavelength information
  Appears idealized. Is this measured information or modeled?
  Out of band transmission information not provided.
  Bandwidth information is a single number; information not provided on the shape of the bandpass.

## Slide 4: Review of Metadata Files
BadPixel Map
  Clear
BondJoint Map
  Not sure what to do with this information other than discarding the affected data.

## Slide 5: Review of Metadata Files
Pointing information
  Does not appear to contain valid geometric information on LEISA’s FOV. Both the numbers and the ‘perfectness’ are suspect.
+- 10 mrad (if those are the units)
????

## Slide 6: Data Useability
Simple to read in.
Motion of Spacecraft and nature of instrument requires use of Spice data, etc. for detailed analysis, which I did not perform.
Although the data is calibrated, significatn work (time varying NUCs, and corrections for s/c pointing) would need to be done to make the data useful. This is a typical ‘feature’ of IR spectral imaging data.

## Slide 7: Review of Data File
Chose the high resolution image from July 14
  Data file only contains radiance data
    Data quality is consistent with MCT detectors. Much residual detector non uniformity, secular variations, and apparent ‘popcorn noise’. Insufficient time to characterize in detail.
    Radiance appears to be higher than expected (compared to Thekaekara, 1973)
Briefly looked at more distant obsevations

## Slide 8
Line number & Wavelength Position
Column Number
Resolution = 240
Resolution = 560

## Slide 9
Wavelength (microns)
Radiance (ergs/cm2/s/A/Sr

## Slide 10
Wavelength (microns)
Radiance (ergs/cm2/s/A/Sr

## Slide 11: Additional Information I would find of use if analyzing the data
For each cube, provide an accompanying cube of:
  Darks used in the calibration
  Flat field used in the calibration
  Gain used in the calibration
  Offset used in the calibration
For each line on the detector, provide:
  A filter transmission spectrum to aid in developing custom spectral deconvolution procedures