Joshua Kammer
Rosetta Alice
End of Mission (EOM)
PDS/PSA Science Archive Review

---
Rosetta Alice
- Ultraviolet Imaging Spectrograph
  - Bandpass: 700-2050 Å
  - Spectral Resolution:
    - 5.6 Å (point source)
    - 12.5 Å (extended source at 700 Å)
    - 9.8 Å (extended source at 2050 Å)
  - Active FOV: 0.1° x 6.0°
  - Nominal Effective Area:
    - 0.03 cm2 (at 1900 Å)
    - 0.53 cm2 (at 1150 Å)

---
R-Alice Datasets
- Spotchecked data from entire mission
  - Prelanding (PRL V2.0)
  - Escort (ESC1 V2.0, ESC2, ESC3, ESC4 V1.0)
  - Extension (EXT1, EXT2, EXT3 V1.0)
- Three data processing levels
  - Level 2: raw data
  - Level 3: calibrated data
  - Level 4: high level / resampled data
- 24 datasets total
- Review Conclusions
  - Data are ready to be released
  - Only minor/editorial RIDs (001-004)

---
Evaluation Tools
- Dell Inspiron i7559 Laptop
  - Ubuntu 16.04 LTS
    - Python 3.6.2
    - Pandas 0.20.3
      - pandas.read_table to read ASCII data
    - Astropy 2.0.2
      - astropy.io.fits to read all FITS files
  - Windows 10 Pro
    - NASAView 3.15.0
      - To check ability to read PDS labels

---
Datasets Review
- Calibration process
  - Well-described (/document/calibration_cookbook.pdf)
  - calib/ directory contains all needed files for utilizing effective area, wavelength vectors, and dark counts at multiple voltage settings (-4000, -4100, -4200, -4300 V)

---
Datasets Review
- Attempted replication of published results
  - Keeney et al. 2017, MNRAS, 469, S158-S177
  - UV starlight (HD 66006) absorption by cometary coma
  - Appulse: 2015 Aug 10 04:28:49, ~21 min
    - Mission phase ESC3
  - Revisit: 2016 Jun 6 03:27:48, ~30 min
    - Mission phase EXT2

---
Datasets Review
Keeney et al. 2017, Fig. A10

---
Datasets Review
Keeney et al. 2017, Fig. A10

---
Datasets Review

---
Datasets Review

---
Datasets Review
My attempt to replicate
Keeney et al. 2017, Fig. A10

---
Review Summary
- Archived data appeared complete and easily interpretable
- RIDs: 4 minor
  - 001: EAICD content
  - 002: Level 4 BROWSE images
  - 003: Additional EAICD text out of date
  - 004: User manual missing information

---
RID 001
- Classification: Minor
- Location: /document/eaicd_alice_v4.pdf
- Title: EAICD content
- Description: The information on CODMAC level 4 data in sections 4.3 and 4.6 are copied from the level 3 sections (4.2 and 4.5) and still reference level 3 data. Sections 4.3 and 4.6 are not listed correctly in the table of contents.
- Recommended Solution: Update these sections of the EAICD to reference level 4 information, and correct the table of contents

---
RID 002
- Classification: Minor
- Location: /browse/*
- Title: Level 4 browse images
- Description: The browse images display the data in units of photons cm-2 s-1 Angstrom-1 sr-1, while the level 4 data is in units of photons cm-2 s-1 Angstrom-1, making direct comparison difficult
- Recommended Solution: Update the browse images to use the same units as the *LIN.FIT files to which they correspond (photons cm-2 s-1 Angstrom-1)

---
RID 003
- Classification: Minor
- Location: /document/eaicd_alice_v4.pdf
- Title: Additional EAICD text out of date
- Description: Section 3.1.3 (Data Directory Naming Convention) does not mention the level 4 data. Section 3.4.3.5 indicates that the BROWSE directory is not used, but it is. Section 3.4.3.12 describes the DATA directory, but also does not mention the Level 4 data. Typo in section 2.3.4 (CODEMAC->CODMAC)
- Recommended Solution: Update the text to include reference to Level 4 data and what is included in the BROWSE directory. Fix minor typo (CODEMAC->CODMAC)

---
RID 004
- Classification: Minor
- Location: /document/alice_manual_v2_1.pdf
- Title: User manual missing info
- Description: Attachments 1-10 at the end of the manual are missing/blank
- Recommended Solution: Include the relevant attachments, or remove the blank pages if the attachments are unnecessary