## Slide 1: PDS review of ROSINA
Mark Perry
Johns Hopkins Applied Physics Lab
9 October2017

## Slide 2: Summary of findings
Despite the good work performed by the ROSINA team, the dataset is still insufficient for analysis and does not meet PDS standards.
  Data set is extensive and complex.
  Documentation was likely prepared by someone familiar with the data, so some crucial information—obvious to the team—is not included.
Primary deficiencies
  Need additional documentation to understand the data.
  Calibration and many essential parameters undefined.
  A users guide would fill most of the gaps (not a PDS requirement?).
These are great instruments and great data; the team should invest the time to make their data available and usable.

## Slide 3: List of RIDs
Insufficient documentation (major)
Label parameters undefined (major)
Data missing from labels (major)
File organization and description (major)
Missing calibration information (major)
Need a data analysis guide (major)
NASAVIEW does not read these data files (minor)

## Slide 4: Insufficient documentation
The dataset does not contain all documentation needed to use and understand its data without prior knowledge.
  Please see other RIDs for suggestions relevant to individual documents.
Suggest that the ROSINA team engage a non-ROSINA scientist to analyze ROSINA data using the PDS documentation and datasets.
  This scientist then produce an updated set of documentation that includes necessary information that is not yet in the PDS documentation but was obtained by the scientist from the ROSINA team during his or her study of the data.

## Slide 5: Label parameters undefined
Most of the parameters in the data-file embedded labels are undefined.
  Hundreds of parameters in the label headers are undefined. None of the PDS documentation contains either definitions or descriptions of them.
  Without these definitions and descriptions, initial work with the dataset is frustrating.
The ROSINA ICD would be a good place to define and describe each of the parameters in the labels of the data files.
  Often, this is done with a table in the appendix.
  Descriptions (not just definitions) facilitate use.

## Slide 6: Data missing from labels (major)
The embedded labels for DFMS and RTOF do not contain information such as
  1) whether the measurement was of ions or neutrals; and
  2)whether the data were taken with high or low resolution.
Add this information to the labels or to the file names.
  This information might be contained in the "mode" parameter, but that parameter is not defined in any documentation (see previous RID)

## Slide 7: File organization and description (major)
Finding a file with the desired information requires exploration and time-consuming examination of the data files.
  For example, each DFMS files are organized by amu.
    This method of organizing the data is not described in any of the documentation.
    There is no way to easily identify which files cover a specific mass.
  Another example: the COPS files cover 5 minutes, and that is not described.
  Another example: the files are not organized by high or low resolution, and there is no easy way to find data with a specific resolution.
There are many ways to correct these deficiencies. For example:
  Information on file organization could be included in DATASET.CAT.
  For DFMS, the central mass could be in the file name.
  Files could be organized in separate folders for high and low resolution.

## Slide 8: Missing calibration information (major)
There are many deficiencies in the calibration data.
  A rough calibration is necessary
    For interpreting the data and
    As a initial point for deriving a higher-fidelity calibration.
    Even a rough calibration provides science; for example, it can permit identification of major species and relative abundances.
  The equations for mass calibration are provided in the ICD, but the calibration parameters are only provided for DFMS CE.
    Also, the equation for the RTOF has two typos.
  There are no data for converting counts (or events) to physical units.
    There is not even a discussion
    There is no information that can be used to begin a conversion: no efficiency or geometric area data or integration times, etc.
  There are no data on fragmentation.
  The calib directory is empty.
It is necessary to provide calibration information, at least for a rough calibration of both the mass scale and the conversion to physical units.
  Provide ranges if exact values are unavailable or inappropriate. Provide caveats, if appropriate.

## Slide 9: Need a data analysis guide
These are complicated data, and a user guide is arguably necessary to facilitate analysis and generate a greater interest and use in these extraordinary data.
  Examples of user-guide contents:
    Provide basic descriptions of each data set
    Define each parameter in the data files, and have an example on how to use them.
    Include examples such as reproducing a published plot from PDS data.
  The user manual included in the documentation folder is for operating ROSINA, and has little utility for data analysts.
Recommendation: provide a user guide for the dataset.

## Slide 10: Additional detail follows

## Slide 11: ICD Section 1.4, “Intended Readership”
The staff of the archiving authority (Planetary Science Archive, ESA, RSSD, design team) and any potential user of the ROSINA data.
  However, it is not intended that people not familiar with the ROSINA sensors and with mass spectrometry are able, based solely on this document and the archived data, to work with ROSINA raw data. This instrument is by far too complex to be understood by laymen. Raw data depend on too many parameters hidden in the housekeeping data to be of any value to the general public. In order to work with raw data one has to familiarize himself with the complete user manual (including the annexes) and one has to be knowledgeable in the field of mass spectrometry.
These are valid points and concerns (and true for many instruments). However:
  The PDS is not for “laymen” and the the “general public,” but for research scientists.
  The supplied user manual is insufficient, even for experts in mass spectrometry.
    The annexes are not yet included.
  The dataset is unique, valuable, and a great resource to the planetary-science community. It would be a shame if the ROSINA dataset were not fully utilized.

## Slide 12: Parameters in data labels are undefined
"ROSINA_DFMS_SCI_PACKET_COUNT "," ","0 "," "
"ROSINA_DFMS_SCI_MASS "," ","16 "," "
"ROSINA_DFMS_SCI_MODE "," ","222 "," "
"ROSINA_DFMS_SCI_MG_FLAG ","OFF "," "," "
"ROSINA_DFMS_SCI_ISB_FLAG ","OK "," "," "
"ROSINA_DFMS_SCI_CEM_THR_FLAG ","OFF "," "," "
"ROSINA_DFMS_SCI_ESAO_PLATE "," ","314.1 ","V "
"ROSINA_DFMS_SCI_ESAI_PLATE "," ","-326.5 ","V "
"ROSINA_DFMS_SCI_MAG_TEMPERATURE "," ","5.312 ","DegC "
"ROSINA_DFMS_SCI_COVER_POSITION "," ","0 "," "
"ROSINA_DFMS_SCI_GCU1_O_PRESSURE "," ","N/A "," "
"ROSINA_DFMS_SCI_MOTOR_POSITION "," ","0 "," "
"ROSINA_DFMS_SCI_GCU2_O_PRESSURE "," ","N/A "," "
"ROSINA_DFMS_SCI_LEDA_FAR_OFFSET "," ","20 "," "
"ROSINA_DFMS_SCI_G_AT_PEAK "," ","0 "," "
"ROSINA_DFMS_SCI_GAIN_AT_PEAK "," ","16 "," "
"ROSINA_DFMS_SCI_P_PT_AT_PEAK "," ","5 "," "
"ROSINA_DFMS_SCI_GAIN_OF_SPECTRUM"," ","16 "," "
"ROSINA_DFMS_CMD_ERROR_FLAG ","OFF "," "," "
"ROSINA_DFMS_HV_SAFE_FLAG ","OFF "," "," "
"ROSINA_DFMS_COV_MOT_DIS_FLAG ","ON "," "," "
"ROSINA_DFMS_HK_COUNTER "," ","26189 "," "
"ROSINA_DFMS_CMD_COUNTER "," ","28689 "," "
"ROSINA_DFMS_CMD_ERROR_CNT "," ","8 "," "
"ROSINA_DFMS_CMD_ERROR_POS "," ","0 "," "
"ROSINA_DFMS_SCI_COUNTER "," ","4573 "," "
"ROSINA_DFMS_SCI_ERROR_COUNTER "," ","22 "," "
"ROSINA_DFMS_SCI_ERROR_POSITION "," ","0 "," "
The data labels are embedded in the data files.
Each data file has 300 to 400 parameters.
  Excellent; probably necessary.
  Only spacecraft target position and velocity vectors are discussed.
In particular, “mode” seems as if it would be useful.
Dangerous to assume an understanding just from the parameter name.
  Every one of these should be defined in the ICD.

## Slide 13: Data values are identified in rtof_data.fmt
OBJECT = COLUMN
NAME = COUNT
DESCRIPTION = "Channelnumber. The values are in the
range from 1 to 131099 and ascending."
UNIT = "CHANNEL_NUMBER"
DATA_TYPE = ASCII_INTEGER
START_BYTE = 1
BYTES = 6
END_OBJECT = COLUMN
OBJECT = COLUMN
NAME = HISTOGRAM
DESCRIPTION = "Histogram data of RTOF ETS. Field
contains 0 for ETSL"
UNIT = "EVENT_NUMBER"
DATA_TYPE = ASCII_INTEGER
START_BYTE = 8
BYTES = 17
END_OBJECT = COLUMN
OBJECT = COLUMN
NAME = EVENT
DESCRIPTION = "RTOF Event data of either ETS or ETSL"
UNIT = "EVENT_NUMBER"
DATA_TYPE = ASCII_INTEGER
START_BYTE = 26
BYTES = 17
These data fields should be described in the ICD.
Can deduce that “Channel number” is the mass field.
Can deduce that “Event number” is related to counts or density.
No explanation of why there are two “Event numbers.”

## Slide 14: Calibration equations (red=typo, green=particularly useful)
2.5.3.2 Mass scale calculation for DFMS CE
  m(stp)= m0-(wdth0*sqrt(m0)/stw) +(stp-1)*m0/stw
  with m0: central mass, corresponds to commanded mass(ROSINA_DFMS_SCI_MASS), but may be sligthly shifted due to temperature effects, shift can be deduced from known masses, especially inflight gas calibration modes
  wdth0: total scan width/2; =140 for LR; = 280 for HR
  stw: =stepwidth; =4000 for LR and 40000 for HR
  stp: step number
2.5.3.4 Mass scale calculation for RTOF
  m(chn)=const*(chn*1.5-t0)^2
  with chn: channel number const and t0 derived from (at least) two known mass peaks (m1 and m2 at channel chn1 and chn2) of the spectrum, temperatur dependent:
  t0=(sqrt(m1/m2)*chn1-chn2)*1.5)/(sqrt(m1/m2)-1)
  const=m1/(chn1*1.5-t0)^2

## Slide 15: A plot of RTOF CEM sensor data
Use coarse mass calibration from ICD.
  Determined that it is likely the low-resolution mode; but not certain.
Vertical axis is “Event number”
  Counts? Conversion to density?
Neutrals or ions?
  Maybe this can be deduced from the undefined parameters provided in the label.
From instrument paper

## Slide 16: DATASET.CAT could contain much more information
Descriptions of the data coverage
  Were data obtained continuously for all instruments, or only at certain times or in certain conditions? What were those times or conditions?
  Are there gaps in coverage due to lost data or because not yet processed?
Descriptions of the data files and how they are organized.
  For example, COPS files each cover 5 minutes and there are pressure readings every 2 seconds; why are there zeros in RTOF data changes 30 to 29000?; DFMC noise floor is 52,000; why is there a difference between LEDA rows A and B?; what mass ranges were commanded for the RTOF measurements?; etc.

## Slide 17: Other missing information
No RTOF FAR data.
No ERRATA.TXT file
RTOF_MASS_CAL.FMT file referenced but not provided.
“Appendix AD1-Instrument modes DFMS” referenced (in ICD Section 2.3.1) but not provided.

## Slide 18: Other thoughts
Most problems are systemic and not adequately addressed by adding a few notes or by fixing only the specific examples provided.
COPS data (pressure) are calibrated and may be ready to use.
  Missing some descriptions of the dataset such as time periods, parameters.
  Did not do thorough review of the COPS documentation.
Could remove some repetition.
  Example: instrument requirements and design are in three places: instrument.cat, ICD, operations manual. And also in the referenced SSR paper.
Worth repeating:
  The dataset is unique, valuable, and a great resource to the planetary-science community. It would be a shame if the dataset were not fully utilized.