## Slide 1: RIDS from PDS review of ROSINA
Mark Perry
Johns Hopkins Applied Physics Lab
10 October2018

## Slide 2: Questions on PRL 3 DFMS CE
All the files say quality 4, which means insufficient data to calibrate.
  Is that correct? Some seem to have good signal.
Mass calibration for the GCU data (mode 160) seems to be low by 0.05 to 0.1 u
  Is that the effect noted in Section 3.9 of the Data User Manual: “Species which don’t contain hydrogen are always left of the integer mass”?
mtp8 higher precision for commanded mass of 18u is in MC, only.
  CE data have two significant figures.

## Slide 3: COPS uncertainty not provided
Comparing COPS densities from RTOF, ESC2, mtp14, Level 5: H2O file and CO2 file.
COPS densities usually within 5% of each other.
In one examined file, have difference factors of two to three (see green ovals).
Uncertainties not provided in either Level 4 or Level 5 files (the Data User Manual says that COPS was designed for “a typical error of 4%-6%”).
Time (days)

## Slide 4: Uncertainties for level-5 data are all 20%
The real uncertainties appear to differ.
Is the 20% explained in any of the documentation?
Are their plans to provide improved estimates of the uncertainty?

## Slide 5: RTOF data (ESC4 SS) quality
Are any of the RTOF data files quality 0 or 1 (good)? Spot check only found quality 4. (The gas-calibration files are quality 0).
How does a user find the higher-quality data? Could that information be included in the errata file or other documentation?

## Slide 6: Fragmentation and sensitivity data
Spot checks of fragmentation patterns are correct.
Referenced file name appears incorrect
  The rtof_mode_id_table_*.tab (in ESC4 calib directory) references FRAG00 and SENS00.
  In the directory, there is only rtof_frag01_*.asc
    Has “01” instead of “00”
    No “sens” in the filename

## Slide 7: Notes on documentation
It takes an effort to get comfortable with the data, but the extensive documentation makes that comfort possible.
  Documentation was and is a huge effort and is much appreciated.
Data user manual
  New, needed, and useful.
  The draft does need another round of editing for cleanup. Some examples:
    Defining acronyms, missing figures, location of important files, additional details such as file names for referenced documents.
The aareadme.txt and dataset.cat files are much improved and great.