## Slide 1: PDS Review Rosetta RPC-IES
Yaxue Dong
CU-Boulder/LASP
University of Maryland, College Park, MD
Oct 8-9, 2018

## Slide 2: Overview
ro-c-rpcies-5-ext2-v1.0
  Rosetta Plasma Consortium – Ion and Electron Sensor
  Extension 2: Apr 6 2016 to Jun 30 2016
  L5 data: Ion and electron moments (density, velocity, and temperature)

## Slide 3
voldesc.cat
moment data
/document/docinfo.txt
Not listed in docinfo.txt

## Slide 4: Read the data
Tools:
  PDS3 IDL software: works well
Data structure for both electron and ions:
  SPACECRAFT EVENT TIME UTC
  MASS
  CHARGE
  DENSITY
  VX
  VY
  VZ
  VELOCITY
  TEMPERATURE
  DURATION
  QUALITY FLAGS

## Slide 5
Ion data
n (cm-3)
Velocity (km/s)
Speed (km/s)
Ti (K)
Duration (s)
Charge (e)
mass (amu)
q = 1, 2 e
m = 1, 4 amu
Identifying and separating H+ and He++ signals explained in “moments_calculation.pdf”

## Slide 6
Electron data
q = -e
m = 0 for electrons
m = 1 amu: negative ions (H-)?
ne >> ni explained as photo-electron density much larger than solar wind electrons (moments_calculation.pdf)
n (cm-3)
Velocity (km/s)
Speed (km/s)
Ti (K)
Duration (s)
mass (amu)
charge (e)

## Slide 7: Negative ions?
In electron moment data
m = 1 amu, q = -e
No descriptions of how to identify these negative ions (H-) from electron
Very low density with repeating values: 0, 0.001, and 0.002 cm-3
When density is 0, there is still non-zero velocity data
n (cm-3)
V (km/s)

## Slide 8
ni (cm-3)
ne (cm-3)
Close to the comet:
Alt ~ 5km
Far away from the comet:
Alt ~200 km
ne >> ni explained as photo-electron density much larger than solar wind electrons (moments_calculation.pdf)
Comparable electron density between far away from the comet and the closest approach. Shouldn’t there be much more photo-electrons at the closest approach
Photo-electrons?
  Can we tell it from the energy spectra?
  How does the electron density compare to total the ion density including heavy ions from ICA?

## Slide 9: Summary
One major RID:
  The negative ion data in electron moment files doesn’t look right.
One minor RID:
  “/document/docinfo.txt”: some documents not listed
Editorial:
  In “voldesc.cat”: counts data  moment data
Others: more explanation needed for photo-electrons

## Slide 12
RPC-MIP
RPC-IES