## Slide 1: Review of Rosetta ROSINA DFMS Time Series Abundances urn:nasa:pds:ro-rosina:dfmi_ts_abundance::1.0
Joe Westlake

## Slide 2: Dataset Overview
This is a set of abundances from the Rosetta dual focusing mass spectrometer during the prelanding phase and the escort phase 1
The dataset contains time-series data at specific masses (16, 18, 27, 28, 30, 32, 44) in two files for each mass (labeled hra and hrb)
The dataset contains relevant spatial data (XYZ, latitude, longitude, local-time) and two derived values – “sum of all peaks” and “area under mass peak”

## Slide 3: Documentation Review
Does the dataset contain all documentation needed to use and understand its data without prior knowledge?
  The dataset is severely lacking in documentation. I assume that this is connected with a publication of some sort giving details on how the fits were produced, calibration parameters, etc…
  The label of the files as hra and hrb is not stated anywhere
Is the provided documentation well organized, clear and self-consistent?
  The provided documentation is insufficient. The only information that is given is a snippet of where the data is taken from gives no details on the experiment.
Can the dataset be understood without any external documentation it references, or should the information in said external references be incorporated into the dataset?
  It seems as though this collection is in support of a publication, but I was unable to find the reference. Adding that external reference would be helpful, or explaining it within the dataset would be helpful
If reviewing calibrated data, does the documentation fully explain the calibration process and contain all necessary parameters needed to repeat it?
  The data appears calibrated and perhaps even derived as it has some Gaussian fit of the peaks and the derived values are given. No calibration details are given in the dataset, but calibrations are provided in the instrument papers. They should be referenced and the authors of this dataset should mention which raw datasets they worked with to produce this derived dataset

## Slide 4: PDS Labels and Meta Data
Are the descriptions and scientific content contained inside the PDS labels sufficient to understand their corresponding data products?
  The descriptions give the minimal information needed to figure out what is in the dataset.
Is all significant meta data included directly in the PDS labels?
  yes
Do the labels provide all essential description of data values directly in the label, instead of deferring them to external references or documentation?
  It would be useful to have a definition of the “Cheops reference frame”
Can the data be read programmatically using only the information contained in the PDS labels?
  yes

## Slide 5: Data
Does the data look physically reasonable when examining it by eye or via a display tool?
  Yes, I was able to plot the data against distance from the comet and found an exponential like falloff, which appears consistent with expectations (see following slides)
When displaying the data as plots or images, are there any unexpected deviations?
  No
Formulate a scientific inquiry and attempt to use the data to answer the inquiry.
  If I look at this dataset you can see the differences in activity of the comet over time
If reviewing both raw and calibrated data, attempt to calibrate a raw data file.
  This is a derived dataset, so no calibration information

## Slide 6: Sample Data Plot – H2O, HRA
Chart: 
The plot gives “sum of all peaks” or a pseudo density versus altitude
The exponential fall-off is expected for a comet with escaping gasses
There are several overlapping data points from the multiple time-periods

## Slide 7: Unclear what the hra/hrb difference is?
There are two data files for each. The HRA and HRB.
They have significant differences
Both of these are H2O or Mass 18, but blue is HRA and orange is HRB
Chart: 

## Slide 8: Summary
The dataset is usable as is, but needs significant additional information to be valuable to the scientific community and to stand on its own.
  The authors should include a reference or explain the process of deriving this data set from the primary calibrated data set
  HRA and HRB should be explained somewhere
  The collection.xml file states that this data is from the “prelaunch” phase when it is actually from the “prelanding” phase
  There are inconsistencies in the description between what is on the webpage and the collection.xml description