## Slide 1: Lucy Dinkinesh Data Review
Yuna Kwon
(Caltech/IPAC)

## Slide 2: Table of Contents
1.1 Dinkinesh Raw Data Collection
1.2 Dinkinesh Calibrated Data Collection
1.3 Calibration Collection
1.4 Document Collection
2.1 Dinkinesh Raw Data Collection
2.2 Dinkinesh Calibrated Data Collection
2.3 Calibration Collection
2.4 Document Collection
Part 1
L’Ralph MVIC
Part 2
TTCam

## Slide 3: #1. data_Dinkinesh_raw
What does the ‘offset’ label (20160) under Array_3D_image mean? (can’t find any description in docs)
I confirmed all data file is properly sorted.
File names follow the product naming rule (2.3.4.1 in SIS).
File labels (dimensions and descriptions) also appear correct.
Data files are correctly opened using the PDS4 viewer and DS9.
L’Ralph MVIC

## Slide 4: #2. data_Dinkinesh_calibrated
What does the ‘offset’ label under Array_3D_image mean? (can’t find any description in docs)
Better to give unit information in the sci images, as in the overview.txt of data_Dinkinesh_raw
I confirmed all data file is properly sorted.
File names follow the product naming rule (2.3.4.1 in SIS)
File labels (dimensions and descriptions) also appear correct.
Data files are correctly opened using the PDS4 viewer and DS9.
L’Ralph MVIC

## Slide 5: #3. calibration
I confirmed the array sizes and dimensions of all the calibrated files and associated space files are identical and properly sorted.
L’Ralph MVIC

## Slide 6: #4. document (1) SIS
The label is correctly loaded via the PDS4 viewer.
Overall, looks good to me. Organized clearly and concisely in data contents and formats.
2.3.2.2 Calibrated Data:
  units (radiance) are written differently from that described in 2.2 (radiance / (counts/sec))
  Better to explain the associated FITS structure (dimensions, etc.) first in the calibration steps.
  Better to add header keywords for TDI records (M4TDIx) in Step 4a.
2.3.4:
  2nd & 3rd lines: redundant expression
  2.3.4.1: missing acronym “SCLK”
FITS Keyword (Page 20): DATE
  Does ISO cal mean ISOT format?
All header keywords are correctly inserted.
L’Ralph MVIC

## Slide 7: #4. document (1) SIS
Minor
  Table of Contents: Page number is a mixture of Romans and Italics.
  Missing acronyms: GSFC (p.7)
  Figure 2-2:
    Remove Figure 3-1 text in the figure
    What does “R/Y” stand for?
L’Ralph MVIC

## Slide 8: #4. document (2) Activities
The label is correctly loaded via the PDS4 viewer.
Descriptions look good to me; for the Dinkinesh flyby, would it be possible to provide the region of saturation (at least approximately)?
L’Ralph MVIC

## Slide 9: #4. document (3) Calibration_Procedure
The label is correctly loaded via the PDS4 viewer.
Page 1 Line 6: Pan’s wavelength range is slightly different from the one specified in Table 2-1 of mvic_sis.pdf
ATsum, XTsum: using more specific keywords regarding “actual” (not “planned”, as the headers have both) summing or at least mentioning this point.
I checked that all the header information required for the calibration steps is correctly provided (including units).
Page 3: How can we get the coefficient information?
L’Ralph MVIC

## Slide 10: #1. data_Dinkinesh_raw
174 FITS files
  I confirmed all data files are correctly loaded in the PDS4 viewer and DS9.
  Data labels appear correct.
  The first five TTCam1 & the five TTCam2: 2000 x 2592
  The rest: 1944 x 2592
Can’t find a strong reason to keep pre-flight and in-flight data in the same folder. Better to keep the test images in a miscellaneous file folder or not provided in this high-level latest version?
TTCam

## Slide 11: #2. data_Dinkinesh_calibrated
174 FITS files
  I confirmed all data files are correctly loaded in the PDS4 viewer and DS9.
  Data labels appear correct.
  The first five TTCam1 & the five TTCam2: 5 products of 2000 x 2592
  The rest: 5 products of 1944 x 2592
TTCam

## Slide 12: #3. calibration
Collection_overview is concisely written and easy to understand.
Two master flat FITS files:
  Each has a 1944 x 2592 master flat & its error image
  I confirmed data labels are well sorted and the files are well loaded in both the PDS4 viewer and DS9.
Input parameter files:
  Better to provide at least a key equation showing the role of each coefficient. Hard to understand without looking into another document.
Bad pixel map files:
  Well loaded both in the PDS4 viewer and DS9.
TTCam

## Slide 13: #4. document (1) SIS
Sect. 1.1: most are the same as the previous sentences.
Typo?: “Fig. 1” in the 5th line of Sect. 2.1.1
Typos: In the document, Zhao & Zhou et al. (2024) are mixed.
Minor
  Table of Contents: Page number is a mixture of Romans and Italics.
  Page 3 Sect. 2.1: TDI used without introducing its full name (and not in the acronym list)
  Missing “IPP” description in the main text and the acronym list
  4th line of Sect. 2.3.2.2.4: different font for “input constants”
  Missing “R/Y” description in the Fig. 2-3 and the acronym list
TTCam

## Slide 14: #4. document (2) User’s guide
IDL code parts skipped: I’m not an expert on this language.
TTCam